## Slide 1: Calibration of Digital Hydrophones
Jay Abel
IQOE Workshop
December 13-14, 2021

## Slide 2
DUT
Projector
Reference
Calibration by Standard Reference
Notes: Most of us are quite familiar with the method of calibrating a hydrophone’s open circuit voltage sensitivity (OCVS) using a projector, a calibrated reference and the hydrophone to be tested. For this method, the hydrophone is only used as a receiver.

## Slide 3
Projector or Transducer
DUT
Transducer
T / R
Calibration by Reciprocity
Notes: In the method of reciprocity, one of three hydrophones must be operated in both transmit and receive modes. By comparing the data obtained from the DUT and Transducer when the projector is sending to the data obtained from the DUT when the Transducer is sending, it is possible to obtain a primary calibration (no reference required).

## Slide 4
Transmit
V
Volume Flow
Electrical Current
Pressure
Voltage
Receive
Notes: The method of free field reciprocity depends on two remarkable facts. First, if a hydrophone is reciprocal (most piezoelectric devices are) then the ratio of volume flow to drive current is identical to the ratio of voltage to pressure. Second, for free field conditions, the ratio of pressure at the receiver to volume flow at the transmitter is a well-known relationship that depends only on the density of the medium and the frequency of interest.
If we know the sending current and received voltage, we can determine the product of the sensitivity of both devices. If a third device is present, we can determine the ratio of the two sensitivities, and can then determine the exact sensitivity of each device, allowing the transducer to be used as a calibrated reference. Finally, by comparing the DUT to the Transducer (in receive mode with the projector running) we can now determine the sensitivity of the hydrophone.

## Slide 5
Transmit
Electrical Current
What if we could make a digital transceiver?
DAC + AMP
AMP + ADC
Transmit:
Tone Generator
Output current signal
V
Notes: If we could make a digital transceiver, the method of reciprocity could be carried out with a pair of digital transceivers and a hydrophone to be tested. But what would a digital transceiver look like? Here is a functional diagram in transmit mode.

## Slide 6
Receive
Voltage
What if we could make a digital transceiver?
AMP + ADC
Receive:
Tone generator off
Output voltage signal
Notes: In receive mode, a digital transceiver is just a digital hydrophone

## Slide 7
10 M
1k
TX_EN
TX Signal from Tone Generator
RX Signal to Amp / ADC
Simplified Circuit
TX_EN | TX Signal | RX Signal
0 | 0V (off, short to ground) | Receive Voltage
1 | Test Tone Output | Transmit Current (1V / mA)
Notes: In this simplified circuit, we switch a low resistance (approximately 1k ohm) between the hydrophone output and ground. With the resistor in, we can measure the current through the piezoelectric element while applying a test signal to the normally grounded terminal. With the resistor out and the transmit signal driven to ground, the device behaves like a normal digital hydrophone.

## Slide 8
Digital Transducer
Digital Hydrophone
Digital Transducer
Digital Reciprocity
Notes: In the method of Digital reciprocity, transceivers are switched between projector and hydrophone mode over the course of two trials. Data collected from the devices takes advantage of a common interface and synchronization mechanisms, greatly simplifying the test equipment required.

## Slide 9: Thank You! jabel@sensortechcanada.com
Notes: In the method of Digital reciprocity, transceivers are switched between projector and hydrophone mode over the course of two trials. Data collected from the devices takes advantage of a common interface and synchronization mechanisms, greatly simplifying the test equipment required.